## All Info
Item Description | Serial Number | 101 | 102 | 103 | 104 | 105 | 106
Iridium | IMEI | 300434063921240 | 300434063823800 | 300434063928220 | 300434063826810 |  | 300434063925210
| RockSEVEN S/N | 14116 | 14110 | 14120 | 14122 |  | 14113
| RockSEVEN CODE | WNN-CLM | SYN-CWQ | SNN-RDQ | SHH-FGC |  | SHN-DTF
PAR | S/N | 23 | 22 | 21 | 20 |  | 12
| m (mA/mmolm-2s-1) | 0.01343 | 0.01415 | 0.01365 | 0.01415 |  | 0.01414
PAR0
Topside Temperature | S/N | 101 | 102 | 103 | 104 |  | 106
| A | 0.0012711 | 0.0011818 | 0.00072045 | 0.0012141 |  | 0.001232
| B | 0.00054191 | 0.00053336 | 0.00069312 | 0.00052225 |  | 0.00051645
| C | 1.1402e-06 | 1.2471e-06 | -1.5902e-06 | 1.4397e-06 |  | 1.533e-06
| Calibration Residual | 0.00081761 | -0.00010959 | 0.01373 | 0.00083358 |  | 0.00087129
| Standard Error | 0.00061914 | 8.5293e-05 | 0.0078268 | 0.00044245 |  | 0.00047173
SST Temperature | S/N | 101 | 102 | -- | 104 |  | 106
| A | 0.0011955 | 0.0011856 | 0.0012026 | 0.0012138 |  | 0.0012156
| B | 0.00052877 | 0.00053192 | 0.00052643 | 0.00052264 |  | 0.00052239
| C | 1.3223e-06 | 1.2807e-06 | 1.359e-06 | 1.4182e-06 |  | 1.415e-06
| Calibration Residual | -0.00027966 | 0.00022502 | -- | 0.00087088 |  | 0.0008352
| Standard Error | 0.00018212 | 0.00014195 | -- | 0.0004346 |  | 0.00040873
Fluorometer | S/N | 21NH0688 | -- | -- | 21NH0689 |  | --
| ADC Reading in DI Water (0mg/L Standard) | 4 | -- | -- | 40 |  | --
| ADC Reading in 40mg/L Standard (x1) | 2081 | -- | -- | 1952 |  | --
| ADC Reading in 40mg/L Standard (x10) | 15700 | -- | -- | 19700 |  | --
| Fluor0 | 4 | -- | -- | 40 |  | --
| f (ADC counts per mg/L at x1) | 0.019259 | -- | -- | 0.020921 |  | --
Camera | Focus (Turns CCW) | 3.5 | 3.5 | 3.5 | 3.5 |  | 3.5
| Resolution | 128x160 | 128x160 | 128x160 | 480x640 |  | 480x640
Pre-Deployment Checks | GPS Battery Voltage | 3.05 | 2.98 | 3.05 | 3.06 |  | 2.99
| RTC Battery Voltage | 3.05 | 3.05 | 3.06 | 3.3 |  | 3.11
| Main Battery Voltage | 9.33 | 9.73 | 9.32 | 9.36 |  | 9.31
| RTC Working | y | y | y | y |  | y
| SD Card Working | y | y | y | y |  | y
| PAR Working | y | y | y | y |  | y
| Topside Temp Working | y | y | y | y |  | y
| SST Working | y | y | y | y |  | y
| Pressure Working | y | y | y | y |  | y
| Tilt Sensor Working | y | y | y | y |  | y
| Fluorometer Working | y | -- | -- | y |  | --
| GPS Working | y | y | y | y |  | y
| Iridium Working | y | y | y | y |  | y
| Camera Working | y | y | y | y |  | y
Deployment Settings | Unit Start Date and Time | 8/7/2018 12:00:00 AM ADT | 8/7/2018 12:00:00 AM ADT | 8/7/2018 12:00:00 AM ADT | 9/22/2018 12:00:00 AM ADT |  | 9/22/2018 12:00:00 AM ADT
| Scheduled Release Date and Time | 4/30/2019 3:00:00 PM ADT | 4/30/2019 3:00:00 PM ADT | 1/9/2019 3:00:00 PM ADT | 3/10/2019 3:00:00 PM ADT |  | 3/1/2019 3:00:00 PM ADT
| Release Type | DBV Technologies | DBV Technologies | Sub Sea Sonics | Sub Sea Sonics |  | Sub Sea Sonics
Burn Wire (Sub Sea Sonics Only) | Time When Programmed | -- | -- | 7/23/2018 08:55:00 ADT
| Burn Wire Voltage | -- | -- | 5.11
| Days | -- | -- | 170
| Hours | -- | -- | 6
| Mins | -- | -- | 8
| Condition | -- | -- | HIGH
Deployment | GOAL Deployment LOC | C2 | C10 / C11 | C12 | East of M5 |  | East of M5
| Deployment Date and Time | 8/13/2018 22:38:00 ADT | 8/11/2018 19:55:00 ADT | 8/10/2018 15:22:00 ADT
| Depth | 44.2m | 49.1m | 61.2m
| Latitude | 71° 13.37' N | 70° 00.915' N | 67° 64.522' N
| Longitude | 164° 15.678' W | 166° 50.413' W | 168° 11.313' W
Note on F17: Author: Temp Sensor Broke after Calibrating. Average of SN 101, 102, 104, and 106 used for these values

## HE18-01 Deployment Sheet
| Serial Number | 101 | 102 | 103
Pre-Deployment | Deployment Site | C2 | C10 / C11 | C12
| Vacuum Start (psi) | 8 | 7 | 6
| Vacuum End (psi) | 10 | 10 | 10
| Vacuum Cap Plug Replaced? | y | y | y
Deployment | Deployment Site | C2 | C10 / C11 | C12
| Deployment Date and Time | 8/13/2018 22:38:00 ADT | 8/11/2018 19:55:00 ADT | 8/10/2018 15:22:00 ADT
| Depth | 44.2m | 49.1m | 61.2m
| Latitude | 71° 13.37' N | 70° 00.915' N | 67° 64.522' N
| Longitude | 164° 15.678' W | 166° 50.413' W | 168° 11.313' W

## AQ18-01 Deployment Sheet
| Serial Number | 104 | 106
Prep | Vacuum Start (psi)
Vacuum End (psi)
Vacuum Cap Plug Replaced?
Deployment | Deployment Site | East of M5 | East of M5
Deployment Date and Time
Depth
Latitude
Longitude
Burn Wire Release | Date and Time When Programmed
Time Zone (UTC or Local)
Burn Wire Voltage
Days
Hours
Mins
Condition

## Iridium SBD Modems
IMEI | Pop Up S/N | RockSEVEN S/N | RockSEVEN CODE
300434063823800 | 102 | 14110 | SYN-CWQ
300434063924230 | 105 | 14111 | WHH-DEW
300434063925210 | 106 | 14113 | SHN-DTF
300434063921240 | 101 | 14116 | WNN-CLM
300434063928220 | 103 | 14120 | SNN-RDQ
300434063826810 | 104 | 14122 | SHH-FGC

## Temperature
PopUp S/N | PopUp_101_Topside | PopUp_101_SST | PopUp_102_Topside | PopUp_102_SST | PopUp_103_Topside | PopUp_103_SST | PopUp_104_Topside | PopUp_104_SST | PopUp_106_Topside | PopUp_106_SST
Coeff-A | 0.0012711 | 0.0011955 | 0.0011818 | 0.0011856 | 0.00072045 | 0.0012026 | 0.0012141 | 0.0012138 | 0.001232 | 0.0012156
Coeff-B | 0.00054191 | 0.00052877 | 0.00053336 | 0.00053192 | 0.00069312 | 0.00052643 | 0.00052225 | 0.00052264 | 0.00051645 | 0.00052239
Coeff-C | 1.1402e-06 | 1.3223e-06 | 1.2471e-06 | 1.2807e-06 | -1.5902e-06 | 1.359e-06 | 1.4397e-06 | 1.4182e-06 | 1.533e-06 | 1.415e-06
Cal. Residual | 0.00081761 | -0.00027966 | -0.00010959 | 0.00022502 | 0.01373 | 0.00020955 | 0.00083358 | 0.00087088 | 0.00087129 | 0.0008352
Standard Error | 0.00061914 | 0.00018212 | 8.5293e-05 | 0.00014195 | 0.0078268 | 0.00012385 | 0.00044245 | 0.0004346 | 0.00047173 | 0.00040873
Calibration Date | 2018-07-11 | 2018-07-11 | 2018-07-11 | 2018-07-11 | 2018-07-11 | 2018-07-11 | 2018-08-27 | 2018-08-27 | 2018-08-27 | 2018-08-27
TB File | TB118192 | TB118192 | TB118192 | TB118192 | TB118192 | TB118192 | TB118239 | TB118239 | TB118239 | TB118239
Note on H3: Author: Temp Sensor Broke after Calibrating. Average of SN 101, 102, 104, and 106 used for these values

## PAR
SKYE PAR S/N | Pop-Up S/N | m (mA/mmolm-2s-1) | PAR0 | Comments
7 | -- | 0.01353 |  | Broken, Never Used
8 | 4 | 0.01423
9 |  | 0.01355
10 |  | 0.01446
11 | 6 | 0.01348 |  | Deployed 2017
12 | 106 | 0.01414 |  | Deployed 2018
13 | 1 | 0.0139 |  | Deployed 2016
14 | 2 | 0.01415 |  | Deployed 2017
15 | 5 | 0.01383 |  | Deployed 2017
16 | 3 | 0.01448 |  | Deployed 2017
20 | 104 | 0.01415 |  | Deployed 2018
21 | 103 | 0.01365 |  | Deployed 2018
22 | 102 | 0.01415 |  | Deployed 2018
23 | 101 | 0.01343 |  | Deployed 2018
24 | 105 | 0.01416

## Fluorometer
Cyclops S/N | Pop-Up S/N | ADC Reading in DI Water (0µg/L Standard) | ADC Reading in 40mg/L Standard (x1) | ADC Reading in 40mg/L Standard (x10) | Date of Calibration
21NH0688 | 101 | 4 | 2081 | 15700 | 2018-07-18
21NH0689 | 104 | 40 | 1952 | 19700